**LH8800 Ventilator**



**specific**

|  |  |
| --- | --- |
| **Model** | **LH8800** |
| Ddiplay | 10.4”LCD screen |
| Ventilation Mode | VCV/IPPV、PCV、SIMV、PSV、CPAP、BIPAP、SIGH、PEEP |
| Gas Source | O2: 280Kpa-500Kpa |
| Air:280Kpa-500Kpa |
| Power Source | AC：100V-240V 50/60Hz |
| Features | * Gas - Electricity separation design ensures safe operation; * Wide application to adults, children; * Exhalation valve can be disassembled and disinfected, which avoids cross-infection; * Seamless switching of invasive and non-invasive ventilation mode; * High integrated modular design of the gas circuit greatly reduces gas leakage , gas loss of system dead space and system compliance and responds rapidly; * Proportional valve with high speed and high precision sends air quickly and accurately, which improves the man-machine synchronization, reduces the consumption of additional respiratory power, provides patients with comfortable excellent breathing; * Intelligent identification of off tube, blocking and other conditions ensures the safety of patients; * Optional: humidifier, medical air compressor. |
| Tidal Volume | 20ml～2000ml |
| Volume/min | ≥18 L/min |
| Breath Time | 0.3-6s |
| Pressure control | 1—50cmH2O |
| I:E | 10:1-1:10 |
| Breathing Rate | 0.5-100Times/min |
| Inspiratory trigger mode | Pressure & Flow |
| Inspiratory pressure trigger | -20 cmH2O～0 cmH2O |
| Flow Trigger | 0～20L/min |
| Maximum safe pressure | ＜125 cmH2O |
| Inspiratory rise time | 0.1—2.0s |
| Sign Pressure | 0～25 cmH2O |
| Pressure Range | 10～60 cmH2O |
| PEEP | 0-25cmH2O |
| Asphyxiation backup ventilation time | 10-60s |
| Oxygen adjusting range | 21%-100% |
| Monitoring Function | VT, MV, FREQ, FiO2%, Paw, CL , Flow[L/min], Pa[kPa]，waveform |
| Alarm Function | * Upper and lower airway pressure limit alarm * Gas pressure shortage alarm * Upper and lower minute ventilation volume limit alarm * Upper and lower tidal volume limit alarm * Respiratory rate alarm * Oxygen concentration alarm * Low battery alarm * Asphyxiation alarm |
| Battery | Built-in high capacity battery can ensure the machine working continuously not less than 8 hours |